## Slide 1
Laue Crystals:
Flat and bent
Manuel Sanchez del Rio
srio@esrf.eu
Advanced Analysis & Precision Unit, MEG/ISDD, ESRF

## Slide 2: Bragg & Laue crystal geometries
R
BRAGG
R
T
LAUE
T

## Slide 3: Bibliography
Darwin (1914)The Theory of X-ray Reflexion Philosophical Magazine 27, 315
Ewald (1917) Zur Begründung der Kristalloptik, Teil III Die Kristalloptik der Röntgenstrahlen Ann. Phys. 54, 519
von Laue (1915) Die Theorie der Röntgenstrahlinterferenzen in neuer Form Encyclopaedie d. math Wiss 5,3, (1931) Ergebn. Exakt. Naturwiss.10133.
Zachariasen (1945) Theory of X-ray diffraction in crystals
James (1950) The optical principles of diffraction by X-rays
Taupin (1961) Théorie dynamique de la diffraction des rayons X par des cristaux déformés Bull Soc Fr. Mineral. 84 51
Takagi (1962) Dynamical theory of diffraction applicable to crystals with any kind of small distortion Acta Crystal 15 1311
Recent books:
Pinsker Dynamical Scattering of X rays in Crystals (1978).
Authier Dynamical Theory of X-Ray Diffraction (2001)
Shvyd’ko X-Ray Optics—High-Energy-Resolution Applications (2004)
crystalpy Guigay Sanchez del Rio JSR (2024)

## Slide 4: DYNAMICAL THEORY (1914) – the darwin width
DARWIN TREATMENT OF DYNAMICAL THEORY (1914) – the darwin width
20160912 Spectroscopy day-out srio@esrf.eu

## Slide 5: Zachariasen Formalism (undistorted)
BRAGG or reflection
LAUE or transmission

## Slide 6: Guigay & Sanchez del Rio (undistorted)
Guigay and Sanchez del Rio (2024) https://doi.org/10.1107/S160057752400924X
R
BRAGG
R
T
LAUE

## Slide 7: Geometrical model (surface model, no bulk effect)
20160912 Spectroscopy day-out srio@esrf.eu
Plane waves (incident and reflected)
Common intersection point on surface (no depth)
Symmetric Bragg is non-dispersive (specular). All other cases are dispersive (non-specular, like gratings or prisms)

## Slide 8: SHADOW4 Bragg only: flat and “bent but locally undistorted”
R
BRAGG
R
LAUE
SHADOW4
SHADOW3
Limited applications

## Slide 9: Distorted crystal (reflectivities)
Takagi, Acta Crystal 15 (1962) 1311
Taupin, Bull Soc Fr. Mineral. 84 (1961) 51
Analytical solutions of Takagi-Taupin (TT) only fow a few cases
Numerical integration of the PDF (2D)
    Regular grids (Finite different methods) Authier, Epelboin, Gronkowski, etc etc
    Finite Element Methods Ari-Pekka Honkanen
If deformation u(z) (e.g., uniform curvature) the PDEs reduce to a single ODE
See also: https://doi.org/10.1364/OE.409961
M. Krisch, Ph. Thesis (1993)
Hölzer et al. Cryst. Res. Technol. 33, 555 (1998)

## Slide 10: Approximated methods for bent crystals reflectivity Sanchez del Rio et al. (2015) https://doi.org/10.1107/S1600576715002782
R
BRAGG
R
LAUE
Multilamellar
Multilamellar
Penning-Polder

## Slide 11: EFFECT OF CRYSTAL ANISOTROPY in bent crystals
M. Sanchez del Rio, N. Perez-Bocanegra, X. Shi & V. Honkimaki, J. Appl. Cryst. (2015). 48, 477-491
111
11-2 or 1-10

## Slide 12: Focusing (Lens equation: monochromatic focus by surface) Guigay Sanchez del Rio https://doi.org/10.1107/S1600577521012480
R
BRAGG
R
LAUE
Not working
Bulk effect (Borrmann triangle)

## Slide 13: Focusing Borrmann triangle (monochromatic) Guigay Sanchez del Rio https://doi.org/10.1107/S1600577521012480
R
LAUE
250 mm-thick flat Si111
8.3 keV
17 keV
Plane incident waves gives non-plane reflected waves
From the quadratic phase we can obtain the focal position
Not possible to directly implement in ray tracing

## Slide 14: Focusing Borrmann triangle Guigay Sanchez del Rio https://doi.org/10.1107/S1600577521012480
R
LAUE
250 mm-thick symmetric bent Si111
8.3 keV
17 keV
Spherical Plane incident waves gives non-plane reflected waves
New lens equation
Only for symmetric Laue

## Slide 15: Laue: polychromatic focusing effects
Polychromatic focusing with flat Laue crystals is fake focusing
Lateral Bragg angle variation
Suppose scattering on surface (no Borrmann considerations)
https://doi.org/10.1016/0168-9002(94)91905-4
http://dx.doi.org/10.1063/1.1146141

## Slide 16: Laue: polychromatic focusing effects: Magic condition
https://doi.org/10.1107/S1600576720016428
https://doi.org/10.1117/12.2525449
Peng Qi, Xianbo Shi, Nazanin Samadi, Dean Chapman
See also: https://doi.org/10.1107/S0021889804023829

## Slide 17: Laue-Laue monochromators
Beijing SRF: Wang et al. (2025) XOPPY+XRT https://doi.org/10.1107/S1600576724011841
ESRF
ID11 future cryogenic LL monochromator
Based on existing ID31 and ID15 monochromators
ID31 XOPPY reflectivity simulations (Penning-Polder)

## Slide 18: More slides

## Slide 19
The fluorescence signal is directly proportional to the internal wave field intensity at the atom position . The position of the fluorescing atom with respect to the diffraction plane can be determined.
Batterman and Cole. Rev. Mod. Phys. 36 (1964) 681
X-ray standing waves
X-ray topography
Simulation of dislocations, stacking faults, cracks, etc.
Y. Epelboin, Progress in Crystal Growth and Characterization, 14, (1987 )465
Grazing incidence diffraction
The angles involved are close to the critical angle. Fresnel specular theory predicts an evanescent wave, which must be treated by the dynamical theory.
S. Stepanov
http://x-server.gmca.aps.anl.gov/
R
BRAGG
R
T
LAUE
Takagi, Acta Crystal 15 (1962) 1311
Taupin, Bull Soc Fr. Mineral. 84 (1961) 51
PROJECT 2: STUDY OF CRYSTAL DIFFRACTION BY FEA
Notes: 19

## Slide 20: Using Dynamical theory in x-ray INSTRUMENTS
Chart: I/Io
Bravin et al.
J. Appl. Phys. 97, 113505 (2005)
Thermally deformed
monochromator
Mocella et al. Nucl. Inst. Meth. in Phys. Res. A 467, 414 (2001).
Mocella et al. J. Appl. Cryst. 36, 129 (2003)
Mocella et al. Proc. SPIE 9510 (2015) Honkanen, et al. J. Appl. Cryst. 49, 1284 (2016)
Mocella & Epelboin J. Appl. Cryst. 32, 154, 1999
Touculou et al . J. Appl. Phys. 97, 113505, 2005
X-ray source
m=+1
m=-1
m=0
IDT
SAW
q
B
Surface acoustic waves
Monochromator+Analyzer in Diffraction Enhanced Imaging
Curved spectrometers, monochromators and analyzers
M. Krisch, Ph. Thesis (1993)
Hölzer et al. Cryst. Res. Technol. 33, 555 (1998)
Honkanen et al. J. Synchrotron Rad. 21, 104 (2014)
M. Sanchez del Rio et al. J. Synchrotron Rad.48, 477 (2015)
Notes: 20

## Slide 21: TAKAGI-TAUPIN EQUATIONS

## Slide 22
4) Extreme asymmetry (backscattering) produces a rainbow effect Shvyd’ko 2006 PRL 97, 235502

## Slide 23
Pc=-sinF if a=45 deg and a=1.
Jones Matrices

## Slide 24: Using Dynamical theory in x-ray TECHNIQUES
The fluorescence signal is directly proportional to the internal wave field intensity at the atom position . The position of the fluorescing atom with respect to the diffraction plane can be determined.
Batterman and Cole. Rev. Mod. Phys. 36 (1964) 681
X-ray standing waves
X-ray topography
Simulation of dislocations, stacking faults, cracks, etc.
Y. Epelboin, Progress in Crystal Growth and Characterization, 14, (1987 )465
Grazing incidence diffraction
The angles involved are close to the critical angle. Fresnel specular theory predicts an evanescent wave, which must be treated by the dynamical theory.
S. Stepanov
http://x-server.gmca.aps.anl.gov/
Notes: 24

## Slide 25: Using Dynamical theory in x-ray INSTRUMENTS
Chart: I/Io
Bravin et al.
J. Appl. Phys. 97, 113505 (2005)
Thermally deformed
monochromator
Mocella et al. Nucl. Inst. Meth. in Phys. Res. A 467, 414 (2001).
Mocella et al. J. Appl. Cryst. 36, 129 (2003)
Mocella et al. Proc. SPIE 9510 (2015) Honkanen, et al. J. Appl. Cryst. 49, 1284 (2016)
Mocella & Epelboin J. Appl. Cryst. 32, 154, 1999
Touculou et al . J. Appl. Phys. 97, 113505, 2005
X-ray source
m=+1
m=-1
m=0
IDT
SAW
q
B
Surface acoustic waves
Monochromator+Analyzer in Diffraction Enhanced Imaging
Curved spectrometers, monochromators and analyzers
M. Krisch, Ph. Thesis (1993)
Hölzer et al. Cryst. Res. Technol. 33, 555 (1998)
Honkanen et al. J. Synchrotron Rad. 21, 104 (2014)
M. Sanchez del Rio et al. J. Synchrotron Rad.48, 477 (2015)
Notes: 25

## Slide 26: OTHER APPLICATIONS
Other devices that exhibit strong dynamic diffraction properties can also be described using dynamical diffraction theory.
Fresnel Zone Plates
Yan et al Phys. Rev. B 76, 115438 (2007)
Multilayers/Superlattices
Bartels et al. Acta Cryst. (1986). A42, 539-545
Osterhoff, Opt. Lett. 37, 3705-3707 (2012).
http://x-server.gmca.aps.anl.gov/GID_sl.html
Dynamical diffraction is also of great importance in the fields of Neutron diffraction and Electron diffraction.

## Slide 27
Some new applications of the Dynamical Theory of Diffraction

## Slide 28: BERRY PHASE induces a kick of the x-ray trajectory
K Sawada, S Murakami, N Nagaosa
Phys. Rev. Lett. 96, 154802 (2006)
Y Kohmura, K Sawada, T Ishikawa
Phys. Rev. Lett. 104, 244801 (2010)
Berry, M.V. Proc. R. Soc. London, Ser. A392, 45 (1984)
Berry, M.V. 1990 Physics Today, 43, 12 34-40

## Slide 29: Time-dependent diffraction
Time-dependence diffraction was studied in the 1990s
“wake” monochromator proposed by G.Geloni
G. Geloni Self-seeded X-ray Free-Electron Lasers
https://pubdb.xfel.eu/record/292720/files/RESUB_self_final%20-%20POST_REFEREE.pdf?subformat=pdfa
Y. Shvyd’ko and R. Lindberg, Spatiotemporal response of crystals in x-ray Bragg diffraction, Phys. Rev. ST Accel. Beams 15, 100702 (2012) dx.doi.org/10.1103/PhysRevSTAB.15.100702

## Slide 30
Final remarks

## Slide 31: WHICH MODEL SHOULD I USE?
1st approximation: dynamical diffraction theory for the perfect crystal
    e.g, flat monochromators, curved monochromators and spectrometers locally “undistorted”, standing waves, etc. XOP, SHADOW
2nd approximation: local applications of the dynamical theory for perfect crystals to distorted ones
    e.g, curved crystals: Multilamellar method (XOP/XCRYSTAL)
3rd approximation: geometrical optics or eikonal theory
    e.g, curved Laue crystals: Penning-Polder method (XOP/XCRYSTAL)
4th approximation: wave optics or Takagi theory
    2D numerical integration: Topography codes adapted to user needs.
    1D numerical integration: curved spectrometers, monochromators, analyzers
5th approximation: wave optics or Takagi theory for many waves (white beams, curved surfaces, incoherent superposition, divergent beams)
Hierarchy of dynamical theories of x-ray diffraction for deformed and perfect crystals
J Härtwig (2001 J. Phys. D 34 A70)
TO DO