## Slide 1: Hard X-ray beamline design and optics consideration
Xianbo Shi
Optics Group, Advanced Photon Source

## Slide 2: Outline
General design consideration
Design example: APS-U ISN beamline
Optics specification
Beamline Optics Design and Simulation Workshop 2025

## Slide 3: Integrated architecture framework for beamline optics design and simulation
Beamline Optics Design and Simulation Workshop 2025

## Slide 4: General beamline design considerations for APS-U
Focal size determined by the source demagnification and the BDA diffraction
Coherence controlled by BDA size
100s nm to micron focusing
Beamline Optics Design and Simulation Workshop 2025

## Slide 5: General beamline design considerations for APS-U
To achieve smaller focal sizes than in case 1, the BDA size in case 2 should be:
  dso — the source-to-OE distance
  dsb — the source-to-BDA distance
E = 10 keV | APS | APS-U
Δ < (μm) | 413 | 23.5
Beam size at 25 m | 949 | 273
BDA transmission | 39% | 8%
Lc, H at 25 m (μm) | 2.3 | 40
Beamline Optics Design and Simulation Workshop 2025

## Slide 6: General beamline design considerations for APS-U
Focal size determined by the source demagnification and the BDA diffraction
Coherence controlled by BDA size
100s nm to micron focusing
BDA size defines the focal size and the coherence
For sub-100 nm focusing
Not recommended for APS-U beamlines
Beamline Optics Design and Simulation Workshop 2025

## Slide 7: APS-U ISN beamline layout
Beamline Optics Design and Simulation Workshop 2025

## Slide 8: Analytical Calculation (Vertical Direction)
Source
U25/U21 revolver
M1
28 m
Flat
Up
= 400 mm
= 2.5 mrad
M2
29 m
Focus
Down
= 400 mm
= 2.5 mrad
= 29 m
= 26 m
BDA V
55 m
Vertical
4 μm
KBV
219.604 m
Focus
Down
= 400 mm
= 2.5 mrad
= 164.604 m
= 0.396 m
Focus at 220 m
Beamline Optics Design and Simulation Workshop 2025

## Slide 9: Analytical Calculation
Source
U25/U21 revolver
RMS size | APS-U Brightness mode
(µm) | 14.8
(µrad) | 2.8
(µm) | 3.7
(µrad) | 1.5
RMS size | APS-U Brightness mode
 | 14.8
 | 2.8
 | 3.7
 | 1.5
Electron beam
Total beam
Photon beam
= 4.6 m
= 0.0496 nm
(25 keV)
Beam size at if undulator center is offset by from electron waist
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
P. Elleaume, Undulators, Wigglers and Their Applications Ch. 3, 69–107, CRC Press (2002).
Beamline Optics Design and Simulation Workshop 2025

## Slide 10: Analytical Calculation (Vertical Direction)
Source
,
BDA V
55 m
M2
29 m
Focus
Down
= 400 mm
= 2.5 mrad
= 29 m
= 26 m
Beam size and divergence before BDA V
= 0.88 for rectangular aperture, 1.22 for circular aperture
( M2 slope error)
Beam size before M2
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
Beamline Optics Design and Simulation Workshop 2025

## Slide 11: Analytical Calculation (Vertical Direction)
Source
,
BDA V size Δ
55 m
M2
29 m
Focus
Down
= 400 mm
= 2.5 mrad
Beam (secondary source) size and divergence after BDA V
Beam size before KBV
= 29 m
= 26 m
= 164.604 m
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
Beamline Optics Design and Simulation Workshop 2025

## Slide 12: Analytical Calculation (Vertical Direction)
Source
,
BDA V size Δ
55 m
M2
29 m
Focus
Down
= 400 mm
= 2.5 mrad
Focal spot size
heigh error determined (wavefront propagation)
FWHM = 2.355
= 29 m
= 26 m
= 164.604 m
= 0.396 m
KBV
219.604 m
Focus
Down
= 400 mm
= 2.5 mrad
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
Beamline Optics Design and Simulation Workshop 2025

## Slide 13: Analytical Calculation - Transmission
Source
,
BDA V size Δ
55 m
M2
29 m
Focus
Down
= 400 mm
= 2.5 mrad
= 29 m
= 26 m
= 164.604 m
= 0.396 m
KBV
219.604 m
Focus
Down
= 400 mm
= 2.5 mrad
M2 transmission
-- mirror reflectivity
BDA transmission
KBV transmission
Total transmission
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
Beamline Optics Design and Simulation Workshop 2025

## Slide 14: Analytical Calculation (EXCEL)
Beamline Optics Design and Simulation Workshop 2025

## Slide 15: Analytical Calculation - CRL
Single lens
Compound refractive lenses (CRL)
Effective aperture due to absorption
Transmission
For Gaussian incident beam with size
can be found numerically as the FWHM of the transmitted beam profile
Transmission is then the integration of the beam profile
X. Shi, et al., Proc. SPIE 10388, 103880C (2017).
Beamline Optics Design and Simulation Workshop 2025

## Slide 16: Analytical Calculation – Two moment bender mirror
  M. R. Howells, et al., Opt. Eng. 39, 2748 (2000).
  L. Rebuffi, L. and X. Shi, Proc. SPIE 11493, 1149303 (2020).
— Young’s modulus of elasticity
Ideal slope
Bender slope
Bender Error
  Peter J. Eng, et al., Proc. SPIE 3449, 145-156 (1998).
Beamline Optics Design and Simulation Workshop 2025
Wednesday 13:00 45’
X-ray mirror system design and specification

## Slide 17: OASYS SHADOWOUI
Beamline Optics Design and Simulation Workshop 2025

## Slide 18: OASYS SHADOWOUI - Results
 | Analytical | Shadow
(μm) | 15.2 | 16.8
(μm) | 5.0 | 6.9
(μrad) | 3.6 | 3.7
(μrad) | 2.8 | 2.8
FWHMfocal, H (nm) | 20.5 | 16.7
FWHMfocal, V (nm) | 19.5 | 15.3
Total transmission | 1.04% | 0.78%
 | Analytical | Shadow
 | 15.2 | 16.8
 | 5.0 | 6.9
 | 3.6 | 3.7
 | 2.8 | 2.8
FWHMfocal, H (nm) | 20.5 | 16.7
FWHMfocal, V (nm) | 19.5 | 15.3
Total transmission | 1.04% | 0.78%
*Mirror reflectivity not considered.
Beamline Optics Design and Simulation Workshop 2025

## Slide 19: Optics specification
Mechanical motion, stability, and misalignment
Optics aberration
Thermal consideration
Beamline Optics Design and Simulation Workshop 2025

## Slide 20: Mechanical motion, stability, and misalignment
, e.g.,
Monochromator relative pitch vibration
Focusing mirror angular vibration
, e.g.,
BDA serves as the new source and isolates the source, FM, and Mono vibration.
Focusing lens positional vibration
Beamline Optics Design and Simulation Workshop 2025

## Slide 21: Mechanical motion, stability, and misalignment
Pitch
Yaw
Roll
Y
X
Z
Beam
KBV pitch angle misalignment using OASYS-ShadowOui
KBV
KBH
= 400
=125
=125.5
= 396
Beamline Optics Design and Simulation Workshop 2025

## Slide 22: Optics Aberration – Mirror Figure Error
Beamline Optics Design and Simulation Workshop 2025

## Slide 23: Optics Aberration – Mirror Figure Error
Beamline Optics Design and Simulation Workshop 2025

## Slide 24: Optics Aberration – Mirror Figure Error
Beamline Optics Design and Simulation Workshop 2025

## Slide 25: Optics Aberration – CRL Thickness Error
Beamline Optics Design and Simulation Workshop 2025

## Slide 26: Thermal consideration
No thermal bump
With thermal bump
Bending corrected
M2 absorbed power
Ray-tracing
FEA
ANSYS
COMSOL
XOPPY ShadowOui loop
SRCalc
Beamline Optics Design and Simulation Workshop 2025