## Slide 1: Curved Crystal Analyzers
Xianbo Shi
Optics Group, Advanced Photon Source

## Slide 2: X-ray Emission Spectrometer Geometry
Non-dispersive (scanning)
  Point-to-point/line focusing
    Spherically bent crystal
    Johann geometry
    Johansson geometry
    Other geometry
Dispersive
  von Hamos
  “In-circle” dispersive geometry
  Flat crystal (miniXES)
Beamline Optics Design and Simulation Workshop 2025
Analyzer expensive
  Detector expensive

## Slide 3: Non-dispersive Geometry
Beamline Optics Design and Simulation Workshop 2025
Resolution and solid angle depends on
  Intrinsic broadening due to bending (XOP Bent_Crystal)
  Geometric effect (ray tracing)
P. Jalal and A. Golamreza, J. Opt. Soc. Am. A 29 (2012) 68-77
Case | R1 | R’1 | R2 | R’2
Johann | 1 | 1 | ∞ | ∞
Johansson | 1/2 | 1 | ∞ | ∞
Spherical Plate | 1 | 1 | 1 | 1
Wittry | 1/2 | 1 | 1 | 1
General focusing | 1/2 | 1 | sin2θB | sin2θB

## Slide 4: Effective Solid Angle
Beamline Optics Design and Simulation Workshop 2025
Case | R1 | R’1 | R2 | R’2
Johann | 1 | 1 | ∞ | ∞
Johansson | 1/2 | 1 | ∞ | ∞
Spherical Plate | 1 | 1 | 1 | 1
Wittry | 1/2 | 1 | 1 | 1
General focusing | 1/2 | 1 | sin2θB | sin2θB
The angle deviation Δθ from the Bragg angle at a given point P(x, y, z) on the crystal surface is expressed in an analytical expression
The energy resolution for a given analyzer crystal is calculated by,
where Δθmax and Δθmin are the maximum and minimum values of Δθ across the whole crystal surface, respectively, ΔθB is the rocking curve width of the analyzer crystal, and θB is the Bragg angle.
a1 | 0
a2 | 
a3 | 
a4 | 
a5 | 
a1 | 0
a2 | 
a3 | 
a4 | 
a5 | 

## Slide 5: Effective Solid Angle
Beamline Optics Design and Simulation Workshop 2025

## Slide 6: Effective Solid Angle
Beamline Optics Design and Simulation Workshop 2025
Effective solid angle for |Δθ|<Δθref/2, Δθref is chosen based on
The intrinsic width of the rocking curve for the bent crystal:
For a Si333 crystal bent to 1 m, the width of the rocking curve is about 2.5×10-5 rad for 6.4 keV X-rays.
The natural width of the emission lines to be measured:
The Fe Kα emission line width is about 1.6 eV (ΔE/E = 2.5×10-4). For θB = 45°, the equivalent angle range is given by (ΔE/E ) tan θB = 2.5×10-4.
The finite source size (sample volume that is illuminated):
A 0.1 mm source at a source-to-crystal distance of 1 m gives an angle range of 1×10-4.
Spectrometer resolution

## Slide 7: Dispersive Spectrometers
Beamline Optics Design and Simulation Workshop 2025
Von Hamos
Flat
Other parallel-detection spectrometers
D. B. Wittry and N. C. Barbi, Microsc. Microanal. 7 (2001) 124–141.
B. A. Mattern, et al, Rev. Sci. Instrum. 83 (2012) 023901

## Slide 8: Crystal anisotropy and biaxial bending
Beamline Optics Design and Simulation Workshop 2025
According to anisotropic elasticity theory, the Hooke’s-law equations give
  Where Sqr are the components of the elastic compliance tensor, u is the displacement vector and is the stress tensor.
For two-dimentional bending with torques M1 and M2, we have:
  Integrating these equations we obtain the displacement vector
And the radii are:

## Slide 9: XOPPY-CRYSTAL
Beamline Optics Design and Simulation Workshop 2025
Multilamellar (ML)
  Bent Bragg
  Bent Laue
Penning-Polder (PP)
  Bent Laue
Symmetric Bragg
Bent Crystal ML
12 keV
Si(333)
Rm = 0.5 m
Asymmetric Laue (χ=35°)
Bent Crystal PP
30 keV
Si(111)
Rm = 20.0 m