## Slide 2
Power transport and management in beamlines
Philipp Brumund and Manuel Sanchez del Rio
brumund@esrf.fr srio@esrf.fr
Advanced Analysis & Precision Unit, MEG/ISDD, ESRF
March 21st 2025

## Slide 3: Outline
1.
Introduction
2.
Source emission
3.
Materials & Physics
4.
Heat Load Components

## Slide 4: Introduction (1) – Power transport and management
Materials + Physical phenomena
Temperature increases
Mechanical stress/strain increase
Power
X-rays (light)
Cooling
Beamline component
Optical performance (“shape/flatness”)
      Mirrors
      Multilayers
      Monochromators
No damage
      Slits
      Filters
      (X-ray optics)
Design Goals
Part 1: Source - Manuel
Part 2: Materials, physics (Manuel/Philipp)
Part 3: Components (Philipp)

## Slide 5: Introduction (2) - Goals
Incoming x-rays
“reflected”
(distorted)
Thermal deformations
No damage
      No melting (temperature)
      Acceptable stresses
power
Scientists don’t want this…
Optical performance (“shape/flatness”)
      Deformation of optics should not distort wavefront
X-ray optics
With initial surface shape

## Slide 6: Introduction (3) – Typical Power Design process
OASYS/
Xoppy
Heat Load
ANSYS (FEA)
Postprocessing
(specific code or OASYS)
Manolo
Achtung !
Philipp

## Slide 7: 2 Source emission
2 Source emission

## Slide 8: Power emitted as synchrotron radiation in a storage ring
EBS: 6 GeV, 0.2 A, R= 844m/(2 ) => P ~ 170 kW.
Heats the vacuum chamber - must be cooled.
EBS Beamlines do not use radiations of bending magnets, that are only used for storing electrons.

## Slide 9: Beamline sources: insertion devices
XOPPY
Source spectrum (Photon “flux” or “spectral power” vs Energy)
top view

## Slide 10: Wiggler (W150 ID17) @ 23.7m
= 0.633 * 6**2 * 1.61**2 * 1.6 * 0.2 = 18.96 kW

## Slide 11: Undulator (CPMU18 2m 6mm gap) @ 23.7m
= 0.633 * 6**2 * 0.88**2 * 2.0 * 0.2 = 7.05 kW

## Slide 12: Undulator emission
XOPPY
U42
Slit
30 m

## Slide 13: Undulator emission
XOPPY
U42
Undulator is tunable

## Slide 14: Undulator emission EBS catalog
Gap
For ID21 U42b:
a=1.9395
b=1.0164

## Slide 15: OASYS: Simulation of source emission
1D: Undulator Flux and Spectral Power (flux/power vs E)
2D: Undulator power density (power vs x, y)
3D: Undulator spectral flux density (flux/power vs x, y, E)
SRCALC Undulator spectral flux density (power vs x, y, Ei)

## Slide 16: Simulation of source emission
Undulator spectral flux density
(flux/power vs x, y, E)
Wiggler spectral flux density
(flux/power vs x, y, E)
y
E
Photon flux
Photon Energy [eV]
1000 5000 10000 15000 20000
x

## Slide 17: 3 Material characteristics
3 Material characteristics

## Slide 18: The Photon-Material interactions: what is “absorbed”?
“Transmitters” (R=0)
1 = A + T
I = R + A + T = 1
“Reflectors” (T=0)
1 = R + A
These coefficients are dependent on the photon-energy
Difficult to calculate ab-initio: use tabulated data, typically from QM calculations
Absorption Attenuation
  Attenuation = what is “removed” from the beam (interesting for end-users)
  Absorption = what remains in the material that is converted to heat (interesting to us)
R, A, T are calculated from refractive index: .
  Linear Attenuation Coefficient: (units L-1)
  Mass Attenuation Coefficient μ/ρ (units L2M-1)
“Attenuation” removes all that “interacts” and is proportional to the total cross section:
  u (= 1.660 540 2 × 10-24 g is the atomic mass unit (1/12 of the mass of an atom of the nuclide 12C),
  A is the relative atomic mass of the target element

## Slide 19: The interaction channels
Interaction “channels”:

## Slide 20: Particle absorption dose
These processes involve a particle “shower”
Electrons are mostly thermalized (absorbed locally in form of heat)
Photons are more mobile, and tend to re-scatter or escape from the material
The energy (dose) deposited (absorbed) in the material depends on the particular geometry.
For very accurate calculations use Monte Carlo Particle Tracking codes
For approximated calculations use the Energy-absorption Mass Coefficient: μen/ρ: μen = μ (1-g) https://www.nist.gov/pml/x-ray-mass-attenuation-coefficients
A pessimistic approach (that we use) is to suppose μen=μ thus overestimating a bit the absorption (stay safe!)
Thermalization: locally deposited energy is transformed in thermal energy

## Slide 21: Practical e.g.: power to be removed - CMPU18 at 6mm gap
x 1000 x e (1.6e-19)
absorb low-energies: attenuators (High-pass filters)
absorb high-energies: mirrors (Low-pass filters)
The integral gives
total power ~1kW
The Front End crops a 7kW beam (at minimum gap)
Compute Flux at 2x1 mm@23m

## Slide 22: Power absorbed by an attenuator
X-ray beam
thickness
material & density
Heat load on filters on flux 2 mm x 1 mm @ 23 m:
U42-EBS
Material | Thickness [mm] | Absorbed Power [W]
Be | 300 | 19
Diamond | 300 | 40
Absorption

## Slide 23: Power Absorbed by Mirror
1st Harmonic
3rd Harmonic
Example of reflection, primarily mirrors @ 30 m
3rd Harmonic
Absorbed power 52 W
Same material, different angles Same angle, different coatings
1st Harmonic
Absorbed power 113 W
Reflectivity
Reflectivity

## Slide 24: Attenuator and Mirror absorption
Diamond (300 mm)
Be (300 mm)
Heat load on mirrors @ 30 m: Si (7 mrad) Diamond (100 mm) + Rh (3 mrad
Heat load on attenuators @ 23 m with a projection of 2 mm x 1 mm:

## Slide 25: Thermal Conductivity
Previous slides: calculation of heat load
  Now: physics to deal with them – (mainly) thermal conductivity
  One-dimensional heat equation [1]
  (neglecting thermal source terms )
  Steady-state () and constant boundary rate of thermal energy and wall size
  (Fourier’s law of heat conduction)
  Thermal gradients inversely proportional to conductivity (for 1-D planar)
[1] A. J. Ghajar and D. Yunus A. Cengel, Heat and Mass Transfer: Fundamentals and Applications. McGraw-Hill Education, 2014.
  : thermal conductivity
  : specific heat
  density

## Slide 26: Typical Conductivities
  Thermal gradients inversely proportional to conductivity (for 1-D planar):
  Thermal conductivity critical to reduce thermal gradients
  Low thermal expansion and high mechanical resistance critical for thermal stresses
Stainless steel
Diamond
Silicon
(LN2)
Coppers
(high )
Quite bad absorbers…

## Slide 27: Other Modes of Heat Transfer
Convection
  Forced convection: internal Water/LN2 cooling with “tubing”
    Cooling linked to convection coefficient from empirical equations (wanted high)
Radiation
  Not discussed here
  At high temperature differences

## Slide 28: Why Finite Element Analysis (FEA)?
Simple geometries: analytical solutions exist
  Continuous solution
Solved: analytical differential equation
Solved: discrete system
Complex geometries: FEA mesh captures nearly any desired shape
  Discrete solution
  FEA easily permits to use non-linear properties
    Conductivities (thermal)
    Stress-strain relationships (mechanical)
  Today: fast to solve & looks better 

## Slide 29: 4 Beamline power components
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
Lenses
4 Beamline power components

## Slide 30: Example in numbers: Power Transport on ID03
Undulator: CPMU16.4 (2m, gap 4mm)
  Total power 20kW
  Science needs power!
  Comparison: French person 2223 kWh/year (4days CPMU)
    And: Thierry Brochard cycling ~50W (rounded) 1xCPMU = 400 Thierrys
Chart: Absorbed Beam Power & Share of Total Power (W)
| 1st Horizontal Slit | Vertical Slit | Combined Slit | Multilayer | Mono
 | 12000 | 920 | 5320 | 1710 | 149
Sample:

## Slide 31: Higher expected future heat-loads
Increased use of cryogenic permanent magnet undulators (CPMUs)
Insertion device
(length)
Gap
(mm)
Total Power
(kW)
Power Density (in
16.1m)
(W/mm2)
Comment
CPMU16.4 (2m)
4.0 mm
20.1
1947
ID03
-
4mm (“mini
-
beta”)
CPMU20.5 (4m)
5.5 mm
29.6
2716
Plan for double CPMU – ID20
Recent front end review
See confluence: https://confluence.esrf.fr/x/e-VwBg
Code: OASYS
Juan Reyes approved!

## Slide 32: Slits: Front end
Jeff Wade
Typical optical components that can be present in a beamline:
Source
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
Horizontal Slit
Glidcop Al-15
18 kW
0.5 kW
4 kW
2 kW
150 W
Absorbed powers
Design goal:
No damage
Vertical plane

## Slide 33: Used: Undulator power density (Power vs x, y)
This OASYS/Xoppy tool is very useful to get the heat load on a slit
undulator
Water cooled beamline slit

## Slide 34: FEA Simulations (e.g. slits): One-way coupled thermo-mechanical FEA
Thermal model
  Cooling coefficient from Dittus-Boelter correlation
  Heat loads projected on surface elements
FEA Software ANSYS

## Slide 35: FEA Simulations Slits: Typical Results
1st Hor. Slit
Glidcop Al-15
Temperature field
Heat load: CPMU20.5 (4m) at 5.5mm gap
Min. principal strains (compression)
Total mechanical strain = elastic + plastic
Usually plastic strains occurring!

## Slide 36: Front End – Horizontal Slit (redesign)
Video of new future design in high-resistance copper CuCr1Zr
Water inlet on one side
Cross section:
  5 channels per side
  Slit cutting to 2mm horizontal
Channels in series
Crossing to other side through hole
Slope optimized for heat load profile
Design conform to specs

## Slide 37: Windows
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
Diamond window
beam
Diamond conductivity
temperature dependence important
18 kW
0.5 kW
4 kW
2 kW
150 W
Absorbed powers

## Slide 38: Power transport on a beamline
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
X-ray beam
Similar to Front end Slits
      But: adjustable
18 kW
0.5 kW
4 kW
2 kW
150 W
Absorbed powers
Filippo Cianciosi

## Slide 39: Multilayer: ID03
Slit 1
Absorption by filters
Absorbed power by multilayer
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors /
multilayer
Monochromator
Focusing system
Sample
KB mirrors
18 kW
0.5 kW
4 kW
2 kW
150 W
Absorbed powers
Work with C. Detlefs
Design goals:
No damage +
optical performance
Design: Carole Clavel, D. Baboulin

## Slide 40: Multilayer: Absorbed power
Wide energy bandwidth
Complex absorption profile on substrate
  Power density profile on surface of x-ray optic
Imported to FEA model

## Slide 41: Multilayer: Thermal deformations
Mirror cross-section can influence deformations
Split in 2 deformation components
  Bending (curvature) with radius (here 73km) can be partially corrected downstream
        With adaptive focusing optics (lenses in case of ID03)
  Residual deformations not
“Optimization” of geometry considering all working energies and parameter sets [1]
[1] P. Brumund, C. Detlefs et al., “Thermal optimization of a high-heat-load double-multilayer monochromator,” J Synchrotron Rad, vol. 28, no. 5, pp. 1423–1436, Sep. 2021, doi: 10.1107/S160057752100758X.
Results for one geometrical parameter set

## Slide 42: Monochromator
Slit 1
Source
Absorption by filters
Absorbed power by monochromator:
All remaining power
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
18 kW
0.5 kW
4 kW
2 kW
150 W
Absorbed powers
Design goals:
No damage +
optical performance

## Slide 43: Power: Primary mirror / Multilayer VS. Monochromators
Both: high power optics
  Important difference: projected power densities on optics surface
  (ID emission)
  Mirrors/multilayers:
    Reflection/layer diffraction
    Small glancing angles
      Low projected power densities:
        Water cooling sufficient
  Silicon monochromators:
    Crystal diffraction with lattice spacing
    Bragg-angle , depending on energy and wavelength , from Bragg-law
      High projected power densities
        Liquid-nitrogen (LN2) cooling necessary
x-rays
x-rays
And: can not be “overfilled”
Very intense hotspot

## Slide 44: Monochromator - Temperature
Mostly made of silicon
non-linearity of silicon thermal conductivity becomes clear

## Slide 45: Monochromator - Deformations
Surface deformations for increasing absorbed powers
  Silicon at liquid nitrogen temperatures
    First concave then convex (with increasing power) – typical silicon cryogenic behaviour [1,2]
      Due to first negative thermal expansion at cryogenic temperatures ()
[1] L. Zhang et al., “Thermal deformation of cryogenically cooled silicon crystals under intense X-ray beams: measurement and finite-element predictions of the surface shape,” Journal of Synchrotron Radiation, vol. 20, no. 4, pp. 567–580, Jul. 2013, doi: 10.1107/S0909049513009436.
[2] P. Brumund, J. Reyes-Herrera, C. Detlefs, C. Morawe, M. Sanchez del Rio, and A. I. Chumakov, “Design simulations of a horizontally deflecting high-heat-load monochromator,” J Synchrotron Rad, vol. 28, no. 1, Jan. 2021, doi: 10.1107/S1600577520014009.

## Slide 46: Lenses - ID03
Typical optical components that can be present in a beamline:
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
Lenses
Front End
Diamond Window
DMM
Filters to pink beam
Pinhole
Ø0.8mm
Al-Lens
CPMU16 (Kmax=2.08)
Generally stack of compound refractive lenses
“Transfocator”

## Slide 47: Lenses - Absorption
Absorbed power in Aluminum lens calculated
  Lens geometry
    , upstream pinhole Ø0.8mm
    Decision for diamond lenses (aluminium above melting temperature)
Absorption in aluminum lens
Applied to FEA model

## Slide 48: 5 Summary
Beamline power simulations usually done in conjunction of
  OASYS heat-load calculations
  FEA thermo-mechanical simulations