## Slide 1: X-ray mirror system design and specification
Xianbo Shi
Optics Group, Advanced Photon Source

## Slide 2: OUTLINE
General introduction to X-ray mirror optics and focusing geometry
Mirror specifications
  Mechanical motion, misalignment, and stability
  Optics aberration
APS-U bender mirrors
  Design concepts
  Additional specs: twist and gravity sag
Beamline mirror operation
  Adaptive mirrors
  In-situ metrology and wavefront sensing
  Advanced control
2

## Slide 3: introduction to X-ray mirrors
Total (internal) reflection
Total reflection
Visible light
Snell’s law
Refractive index of silicon
Total reflection
X-ray
10 keV, Si, = 4.9E-6, = 3.1 mrad

## Slide 4: X-ray mirror shapes
1D focusing/collimating
Elliptical mirror
real source → real focus
Parabolic mirror
parallel beam → real focus
Parabolic mirror
real source → parallel beam
Mirror
Focus
Mirror
Source
Hyperbolic mirror
virtual source
→ real focus
Virtual source
Mirror
Focus
Converging beam
Source
Mirror
Focus
Diverging beam
4

## Slide 5: X-ray mirror shapes
2D focusing
Ellipsoidal mirror
Toroidal mirror only approximates to an ellipsoid near the center of the mirror.
Kirkpatrick-Baez (KB) configuration
Each mirror is a elliptical cylinder
5

## Slide 6: X-ray mirror shapes
Surface shape in mirror coordinate
Source
Elliptical cylinder mirror
Focus
y
z
x
Ellipsoidal:
Elliptical cylinder:
Slope function:
Curvature function:
K. A. Goldberg, J. Synchrotron Rad. 29, 991 (2022). doi:10.1107/S1600577522005793
Exact solution
: real source
: real focus
6

## Slide 7: X-ray mirror shapes
M. Howells, et al., Opt. Eng. 39, 2748 (2000). doi:10.1117/1.1289879
Surface shape in mirror coordinate
Source
Elliptical cylinder mirror
Focus
y
z
x
Taylor expansion
, with
(piston);
(tilt);
(defocus);
(coma);
(spherical);
;
: real source
: real focus
Also works for other shapes
: virtual source
: virtual focus
;
;
;
;
;
;
One-moment
Two-moment
Gravity sag
7

## Slide 8: Focusing geometry
Direct focusing: analytical calculation
Beam size and divergence at focus
= 0.88 for rectangular aperture, 1.22 for circular aperture
Beam size before mirror
Mirror with length , grazing angle , slope error
Note: slope error is not a good criteria for specifying diffraction limited mirrors
Source
,
Focus
X. Shi, et al., Proc. SPIE 10388, 103880C (2017). doi:10.1117/12.2274571
8

## Slide 9: Focusing geometry
Zoom focus
9
X. Shi, et al., Proc. SPIE 10388, 103880C (2017). doi:10.1117/12.2274571

## Slide 10: Focusing geometry
Zoom focus
Source
KB mirror 1
KB mirror 2
WF sensor
Source
Transfocator
KB Mirror
Bender mirror
Hybrid mirror (bender+bimorph)
WF sensor
10

## Slide 11
Mirror specifications
  Mechanical motion, misalignment, and stability
  Optics aberration

## Slide 12: Mirror specifications
APS-U optics ESD example:
Specification of mirror positioners and motors
Specification of mirror stability
Specification of mirror substrate
12

## Slide 13: Mirror specifications
Mechanical motion and misalignment
Pitch angle resolution
Consider from the virtual source position:
= 0.1 means the pitch resolution moves the virtual source by 10% of the source size (RMS)
Equivalent to consider it from the focal position:
= 0.1 means moves the focus by 10% of the focal size.
Source dimension: , , ,
Source
Focus
Pitch
Yaw
Roll
Z
X
Y
Beam
From focal size broadening (less sensitive):
pitch angle change < the angle to increase the virtual source size by 10%
13

## Slide 14: Mirror specifications
Mechanical motion and misalignment
Roll angle resolution
= 0.1 means the roll resolution moves the virtual source by 10% of the source size (RMS) in the perpendicular direction.
In a KB setup, roll resolution of the other mirror:
This is more important for monochromators due to large angle
Yaw angle resolution
Dominated by beam tilting over the coating stripe width, normally in the 0.01° level.
Beam size broadening vs yaw (less sensitive)
Pitch
Yaw
Roll
Z
X
Y
Beam
Source dimension: , , ,
Source
Focus
14

## Slide 15: Mirror specifications
Stability
Pitch
Yaw
Roll
Z
X
Y
Beam
Pitch vibration
Consideration from the source size
For = 0.1,
Same consideration from the focal side
Source dimension: , , ,
Source
Focus
15

## Slide 16: Optics stability specifications
, e.g.,
Monochromator relative pitch vibration
Focusing mirror angular vibration
, e.g.,
BDA serves as the new source and isolates the source and Mono vibration.
Focusing lens positional vibration
X. Shi, et al., Proc. SPIE 10388, 103880C (2017 ). doi:10.1117/12.2274571
16

## Slide 17: Mirror surface aberration
Slope error? or Height error?
Diffraction limited focusing, 1000:1, 70 mm long mirror
s0 – slope of log(PSD) vs log(f)
X. Shi, et al., Proc. SPIE 9687, 968703 (2016). doi:10.1117/12.2241139
For diffraction limited focusing mirrors, the height error is a better measure of the surface quality.
17

## Slide 18: Mirror surface aberration
Slope error? or Height error?
Non-diffraction limited focusing, 100:1, 400 mm long mirror
The on-axis Strehl ratio
with the low frequency slope error:
and the high frequency roughness:
, is the angular size of the image
Church, E. L., & Takacs, P. Z. Opt. Eng., 34, 353 (1995). doi: 10.1117/12.196057
Mirror at 50 m, θ = 3 mrad, E = 10 keV |  | 
 | Source (μm) | W (mm)
APS, H | 640 | 4.5
APS-U, H | 36 | 80
APS-U, V | 15 | 200
18
X. Shi, et al., Proc. SPIE 9687, 968703 (2016). doi:10.1117/12.2241139

## Slide 19: Mirror surface aberration
Simulation
DABAM - an open-source database of x-ray mirrors metrology.
M. Sanchez, et al., J. Synchrotron Rad. 23, 665 (2016).
Mirror figure error specification:
The RMS (PV) height error of the mirror to keep both the focal size broadening and the peak intensity reduction by less than 10%.
L. Rebuffi and X. Shi, Proc. SPIE 11493, 1149303 (2020). doi:10.1117/12.2567501
19

## Slide 20
APS-U bender mirrors
  Design concepts
  Additional specs: twist and gravity sag

## Slide 21: Design concept
Two-moment bender equation (general)
where is Young’s modulus of elasticity, is the mirror length, is the moment of inertia of the beam cross-section, given by
Substituting the exact elliptical equation (radius of curvature):
The variable width is given by
= 60 m
= 500 mm
= 3 mrad
= 1.31e5 N/mm2
= 250 mm
= 10 mm
FEA analysis
L. Zhang, et al., AIP Conference Proceedings 1234, 801 (2010). doi:10.1063/1.3463335
= 0 at the mirror center
21

## Slide 22: Design concept
Two-moment bender equation with trapezoidal/rectangular mirror
M. Howells, et al., Opt. Eng. 39, 2748 (2000). doi:10.1117/1.1289879
L. Rebuffi and X. Shi, Proc. SPIE 11493, 1149303 (2020). doi:10.1117/12.2567501
where is Young’s modulus of elasticity, is the mirror length, is the moment of inertia of the beam cross-section, given by
, with
Rectangular mirror: = 0
22

## Slide 23: Design concept
Two-moment bender equation with trapezoidal/rectangular mirrors
23
p
q

## Slide 24: Design concept
Oasys elliptical mirror bender
Bender vs Ideal profile
Bender error
L. Rebuffi and X. Shi, Proc. SPIE 11493, 1149303 (2020). doi:10.1117/12.2567501
24

## Slide 25: Design concept
Oasys elliptical mirror bender – ray tracing example
L. Rebuffi and X. Shi, Proc. SPIE 11493, 1149303 (2020). doi:10.1117/12.2567501
25

## Slide 26: APS-U bender mirror
APS-U bender design
J. Anton, et al., Proc. SPIE 11100, 111000B (2019). doi:10.1117/12.2529391
Z. Qiao, et al., Proceedings of MEDSI2020, TUOA01, 99 (2021). doi:10.18429/JACoW-MEDSI2020-TUOA01
26

## Slide 27: APS-U bender mirror
TWIST
Before twist correction
After twist correction
Twist correction: originally planned with shimming, finally used the screw hole tolerance
Twist specification <2 µrad/cm length
Z. Qiao, et al., Proceedings of MEDSI2020, TUOA01, 99 (2021). doi:10.18429/JACoW-MEDSI2020-TUOA01
27

## Slide 28: APS-U bender mirror
Gravity sag
,
,
Euler-Bernoulli equation
is the Heaviside step function:
Numerical integration
28

## Slide 29: APS-U bender mirror
Gravity sag
= 300 mm
= 10 mm
= 28.0 mm
= 0.539
= 43.1 mm
= -139
= 139
Source
Focus
29

## Slide 30
Beamline mirror operation
  Adaptive mirrors
  In-situ metrology and wavefront sensing
  Advanced control

## Slide 31: Beamline mirror operation
Adaptive mirrors
JTEC bimorph mirrors
APS mechanical bender mirrors
Hybrid adaptive mirrors
OSAKA-Nagoya-JTEC-APS collaboration
J. Anton, et al., Proc. SPIE 11100, 111000B (2019). doi:10.1117/12.2529391
X. Shi, et al., Proc. SPIE 11491, 1149110 (2020). doi:10.1117/12.2569129
31

## Slide 32: Beamline mirror operation
Adaptive mirrors
APS mechanical bender mirrors
J. Anton, et al., Proc. SPIE 11100, 111000B (2019). doi:10.1117/12.2529391
In-situ metrology (cap sensors)
LTP calibration
APS-U XPCS beamline
bender KB vertical mirror
32

## Slide 33
X-ray Wavefront sensing
Grating interferometry
Speckle tracking
Coded-mask-based Wavefront Sensor
X. Shi, et al., US patent US20220265231 (2022).
Z. Qiao, et al., Appl. Phys Lett. 119, 011105 (2021).
Z. Qiao, et al., Optica 9, 391 (2022).
Z. Qiao, et al., Opt. Express 28, 33053 (2020).
Z. Qiao, et al., Proc. SPIE 11492, 114920O (2020).
Z. Qiao, et al., J. Imaging 7, 249 (2021).
X. Shi, et. al, Proc. SPIE 12240, 122400H (2022).
M. Frith, et al., Rev. Sci. Instrum. 94, 123102 (2023)
Zoomable WF sensor
Compact WF sensor
A standard diagnostic tool optimized for each APS-U beamline.
33

## Slide 34: X-ray Wavefront sensing
Hybrid adaptive mirror testing
Mechanical bender response
Piezo response with no bending
Downstream bent to 40N
Hybrid adaptive mirror
Adjustable Zoom WF sensor
28-ID-B
200 mm
900 mm
Detector
Coded mask
34
Wavefront backpropagation
Focal plane
DO NOT DISTRIBUTE

## Slide 35: AI/ML-driven alignment, stabilization and control of beamline optics
An optics automation system combining adaptive optics, wavefront sensing, and ML/AI control algorithm for beamline auto-alignment/focusing and wavefront control:
Portable and non-invasive
Fully tested on digital twin
User-friendly and configurable via GUI
35
Ultra-realistic digital twin
Zoom Optics
Test &
user training
Autoalignment/autofocusing
Wavefront Correction
Adaptive optics
L. Rebuffi, et al., Opt. Express 31, 21264 (2023). doi:10.1364/OE.488189
L. Rebuffi, et al., Opt. Express 31, 39514 (2023). doi:10.1364/OE.505289

## Slide 36
Acknowledgement
THANK YOU!
  Zhi Qiao (former APS postdoc)
  Matthew Highland
  Matthew Frith
  Luca Rebuffi
  Michael Wojcik
  Lahsen Assoufid
This work was supported by the US Department of Energy, Office of Science, Office of Basic Energy Sciences, under Contract No. DE-AC02-06CH11357.
  Kazuto Yamauchi (OSAKA Uni.)
  Takato Inoue (Nagoya Uni.)
  Yoshio Ichii (JTEC Corporation, now at SLAC)
  JTEC Corporation
  Gautam Gunjala (former UC Berkeley student)
  Antoine Wojdyla (ALS-LBNL)
  Kenneth Goldberg (ALS-LBNL)
  Laura Waller (UC Berkeley)